## Slide 1: Introduction to natural organic matter characterization with ultra-high resolution mass spectrometry
Will Kew
Chemist
william.kew@pnnl.gov
November 7th, 2023

## Slide 2: Overview
What is High Resolution Mass Spectrometry?
Typical Experimental Workflow
Case Studies

## Slide 3: What is Mass Spectrometry?
Mass Spectrometry – measuring the mass of molecules
  Technically, the ‘mass-to-charge’ ratio (m/z) of ions
Molecules are converted to gas phase ions for analysis
  Electrospray Ionization (ESI)
  Matrix Assisted Laser Desorption Ionization (MALDI)
  Electron Ionization (EI)
  Each ionization method has different biases and properties
Mass spectrometers have many types
  Single quadrupoles, triple quadrupoles
  Time-of-flight (TOF)
  Magnetic Sector
  FTMS – Orbitrap or Ion Cyclotron Resonance (ICR)
Liquid sample being infused into an ESI source

## Slide 4: What is High Resolution Mass Spectrometry?
High resolution mass spectrometry (HRMS) resolves isobaric compounds
  Molecules with the same nominal (integer) mass
  e.g. C16H18O10 and C13H22O12
  370.08997 vs 370.11113 – mass difference of just 0.021 Daltons, or 21 mDa
HRMS includes Orbitrap and ICR MS – two types of Fourier transform MS (FTMS)
LC-MS (Orbitrap)
EMSL 21 Tesla FTICR MS

## Slide 5: How High is High Resolution?
7T ICR
Orbitrap
Top shows spectrum acquired with 7T FTICR
Bottom shows spectrum acquired with Orbitrap (Exploris 480)
Sample is Suwanee River Fulvic Acid
Each peak represents a measured exact mass
>20,000 peaks detected in this 7T spectrum
Dozens of peaks at each nominal (integer) m/z value

## Slide 6: Ultra High Resolution with 21 Tesla FTICR MS
C27H22O16, mma 13ppb
C24H26O16S, mma 24ppb
Over 100 peaks at nominal mass 601
Two features resolved despite being only 3.3mDa apart, and formula assigned with ultra high mass accuracy
mma = mass measurement accuracy (ppm or ppb)

## Slide 7: What is High Resolution Mass Spectrometry?
How do you get molecular formula from exact masses?
All isotopes have a unique exact mass
The mass defect is difference from integer (nominal) mass
Combinatorial calculation – with assumptions – allows us to annotate molecular formula to exact masses
e.g. mass of 380.0744 = C17H16O10

## Slide 8: Molecular formulas are not structures!
All of these are C17H16O10
A single chemical formula can be multiple structural or stereoisomers

## Slide 9: Typical Experimental Workflow FTICR MS for NOM Characterization

## Slide 10: Typical Experimental Workflow
Experiment Design
Sites
Sampling and Sample Preparation
Coring
FTICR MS Measurement
Ionization Mode
Raw Data Processing
Data Analysis and Interpretation
Conditions
Hypothesis
Sample Numbers
Replicates
Filtering
Wet or Dry Storage
Anoxic?
Desalting (Solid Phase Extraction SPE)
Solvent Extraction
Direct infusion or online LC?
Instruments
Acquisition Settings (Resolution, Sensitivity, Mass range)
Signal Processing
Peak Picking
Mass Recalibration
Formula Assignment
Spectrum Metrics
Stoichiometric Metrics
Visualization
Peak Intensity Data
Biogeochemical Contextualization

## Slide 11: Typical Experimental Workflow – Experiment Design
Which sites?
  Soil types, agricultural, NEON, geology
What conditions?
  Time series
  Seasonal
  Intervention – Crop rotation, irrigation
  Irregular event – fire, flood, drought
Hypothesis
  “Does drought change the nature of bioavailable soil C?”
  “Does biomass burning change the composition of organic N in soils?
Sample Numbers and Replicates
  ‘Biological’ – i.e. field/condition replicates
  ‘Technical’ – i.e. sample processing/measurement replicate
  MS can have batch, drift, and run-order effects
  Rich data (1000s of features) means statistical power analysis is important!
MONet Sites To Date

## Slide 12: Typical Experimental Workflow – Sampling

## Slide 13: Typical Experimental Workflow – Sample Preparation
Cores are split into sections
Sections are sieved and air-dried
Samples are extracted with water
Extracts are cleaned up by (SPE)
Solid Phase Extraction

## Slide 14: Typical Experimental Workflow – Sample Preparation
For ESI-FTICR MS we need samples in the solution state
  Fully dissolved – no particles!
  Correct concentration
    Not too high – carryover, contamination
    Not too low – poor signal-to-noise
    Typically, 10-50 ppm (mg/L)
MS is sensitive to salts
  Many soils have high mineral content
  Some extraction solvents add to salt level
    e.g. HCl, NaOH extractions
  Desalting can be critical!
Typically, we use solid phase extraction (SPE) to filter, concentrate, and cleanup our samples
Yan Li et al. Analytical Chemistry 2016 88 (13), 6680-6688
DOI: 10.1021/acs.analchem.5b04501
Limnol. Oceanogr.: Methods 6, 2008, 230–235

## Slide 15: Typical Experimental Workflow – FTICR MS Measurement
Typical measurement time – 25 mins
Typical sample usage – 55 µL at 20 mg/L
  Only using ~1 µg per measurement

## Slide 16: Typical Experimental Workflow – Raw Data Processing
Raw Data from Instrument
Automated
Signal processing
Peak picking
Mass recalibration
Formula assignment
Table of peaks, masses, assignments and properties

## Slide 17: Typical Experimental Workflow – Data Analysis and Interpretation
m/z 195.08765
m/z
[C8H10N4O2 + H]+
Stoichiometric Analysis
H/C = 1.25
O/C = 0.25
DBE = 8
AImod = 0
NOSC = 0.75
Caffeine

## Slide 18: Typical Experimental Workflow – Data Analysis and Interpretation
Study of DOM composition of >100 Swedish lakes
Statistical analysis
  Ordination, Correlation
Nat Commun 5, 3804 (2014). https://doi.org/10.1038/ncomms4804

## Slide 19: Recent Review!
Environ. Sci. Technol. 2021, 55, 14, 9637–9656
doi: 10.1021/acs.est.1c01135
Recent review covers key aspects of SOM characterization by FTICR MS experiment design!

## Slide 20: HRMS Strategies for DOM
2020 Book Chapter on DOM analysis by HRMS
Multidimensional Analytical Techniques in Environmental Research. https://doi.org/10.1016/B978-0-12-818896-5.00004-1
Covers state-of-the-art and future directions

## Slide 21: Typical Experimental Workflow
Experiment Design
Sites
Sampling and Sample Preparation
Coring
FTICR MS Measurement
Ionization Mode
Raw Data Processing
Data Analysis and Interpretation
Conditions
Hypothesis
Sample Numbers
Replicates
Filtering
Wet or Dry Storage
Anoxic?
Desalting (Solid Phase Extraction SPE)
Solvent Extraction
Direct infusion or online LC?
Instruments
Acquisition Settings (Resolution, Sensitivity, Mass range)
Signal Processing
Peak Picking
Mass Recalibration
Formula Assignment
Spectrum Metrics
Stoichiometric Metrics
Visualization
Peak Intensity Data
Biogeochemical Contextualization

## Slide 22: Case Studies

## Slide 23: Case Study 1 – “1000 Soils” – the prequel to MONet
WSU1
WSU2
WSU3
Fire events shift soil NOSC values higher

## Slide 24: Case Study 2
Anoxic (sulfidic) vs oxic (non-sulfidic) floodplain sediments
FTICR informed thermodynamic analysis
Oxic conditions = more reduced C
Notes: https://www.nature.com/articles/ngeo2940

## Slide 25: Questions?
william.kew@pnnl.gov

## Slide 26: Thank you